## YL09-example
Y520 Temperature Calibrator | [7 empty] | 34895.91 |  | 11773.31 |  | 12178.98
v0.2.1: 2022-04-07 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards | Distilled water | µS/cm |  | °C | 2.4 | °C | 29.4 | °C | 28.2 | °C
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values (with default K=1 & B=0)
Instructions |  |  |  | mS/cm | °C |  | mS/cm | °C | mS/cm | °C | mS/cm | °C | mS/cm | °C
Calibration Date: | 2022-04-07 | [5 empty] | 0.21351 | -3.9338 | 0.60816 | 20.9072 | 0.60427 | 20.1979
[7 empty] | 0.21351 | -3.9603 | 0.60796 | 20.9132 | 0.60412 | 20.1956
1. GET & paste sensor info from Modbus Runner software. | [6 empty] | 0.21182 | -3.959 | 0.60781 | 20.9706 | 0.60454 | 20.1932
SN: | YL0917101101 | [5 empty] | 0.2135 | -3.9239 | 0.60776 | 20.9766 | 0.60438 | 20.1909
Version: | hw 1.2,sw 1.7 | [5 empty] | 0.21494 | -3.9126 | 0.6077 | 20.9335
Probe: | User Coefficients (previous) | [5 empty] | 0.21494 | -3.9192 | 0.60741 | 20.9383
K: | 1.0386 | (slope) |  |  |  |  | 0.2124 | -3.8709 | 0.60706 | 20.9442
B: | 62.186 | (intercept) |  |  |  |  | 0.21556 | -3.8563 | 0.60747 | 20.9514
[9 empty] | 0.60701 | 20.9598
2. Re-SET "User Coefficients" to defaults. | [8 empty] | 0.60639 | 20.967
K: | 1
B: | 0 |  | Average | #DIV/0! | #DIV/0! |  | 0.21377 | -3.917 | 0.60747 | 20.9462 | 0.60433 | 20.1944 | #DIV/0! | #DIV/0!
|  |  | Std | #DIV/0! | #DIV/0! |  | 0.0013003 | 0.037401 | 0.00052781 | 0.023586 | 0.00017807 | 0.00303 | #DIV/0! | #DIV/0!
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! |  | 0.6083% | -0.9548% | 0.08689% | 0.1126% | 0.02947% | 0.015% | #DIV/0! | #DIV/0!
Type into pale yellow cells G3, J3, M3, P3 |  |  | Converted units: | #DIV/0! | µS/cm |  | 213.7724 | 269.233 | 607.475 | 294.0962 | 604.329 | 293.3444 | #DIV/0! | #DIV/0!
[7 empty] | µS/cm | °K | µS/cm | °K |  | °K |  | °K
4."GET Cond" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y | [9 empty] | X | Y | measured
Paste into pale yellow cells. |  |  |  | Measured | Reference |  | predicted Y | residual Y | calibrated X | [5 empty] | Measured | Reference | Resistence
|  |  |  | 2.4 | 34895.91 |  | 34881.4747 | 14.4353 | 34881.4747 | [5 empty] | -3.917 | 2.4 | 34895.91
5. Delete calibration X:Y pairs that are not valid. |  |  |  | 29.4 | 11773.31 |  | 11462.951 | 310.359 | 11462.951 | [5 empty] | 20.9462 | 29.4 | 11773.31
Delete green values in pale yellow cells D25:E28 |  |  |  | 28.2 | 12178.98 |  | 12503.7743 | -324.7943 | 12503.7743 | [5 empty] | 20.1944 | 28.2 | 12178.98
[9 empty] | 36963.1212 | [5 empty] | #DIV/0! | 0
6. Re-SET "User Coefficients" to new calibration values.
Paste these blue values into Modbus Runner and hit "SET". | [14 empty] | NOTE:
K: | -867.3527 | slope |  | slope | -867.3527 | [9 empty] | Standard practice is for calibration curves to have the known reference values on the X-axis, because they are the "independent" variable.
B: | 36963.1212 | intercept in mS/cm |  | intercept | 36963.1212 | °C | [8 empty] | YosemiTech puts the measured values on the X-axis, despite it being the "independent variable".
Many other well-known sensor companies also follow this practice, even though it does not technically follow statistical best practice.
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.

## Info
[10 empty] | Steinhart-Hart equation |  |  |  |  | https://www.newport.com/medias/sys_master/images/images/h67/hc1/8797049487390/AN04-Thermistor-Calibration-and-Steinhart-Hart.pdf
|  | John's Y520 |  |  |  | Beth's Y520 |  |  |  | 3-term form: | 1/T = C1+ C2 * ln(R) + C3 * ln(R)^3 |  |  |  | https://www.mstarlabs.com/sensors/thermistor-calibration.html
|  | YL0917101101 | [7 empty] | 2-term form: | 1/T = C1’ + C2’ * ln(R) |  |  |  | https://electronics.stackexchange.com/questions/454954/whats-the-simplest-way-to-calibrate-a-thermistor
|  | hw 1.2,sw 1.7 | [12 empty] | https://www.thinksrs.com/downloads/programs/Therm%20Calc/NTCCalibrator/NTCcalculator.htm
| Temp coef. 1 | 1.0386 |  |  | C1' | 1.0264 |  |  | C1'
| Temp coef. 2 | 62.186 |  |  | C2' | 132.3538 |  |  | C2'
[11 empty] | Known Temp C | Measured Temp C | Known Temp K | known 1/T | Resistence Ω | ln(R) | intercept (A) | slope (B) | β (=1/B)
Kitchen Thermometer |  |  |  | 73.1755 | 1/T |  |  | 142.8367 | 1/T | John's Y520 | 0.1 | -6.2018 | 273.25 | 0.0036597 | 39365.63 | 10.5806 | 0.00050937 | 0.00029671 | 3370.2474
0.1 | Temp in Ice water | -6.2018 | *C |  |  | 1.6544 | *C |  |  |  | 2.4 | -3.92 | 275.55 | 0.0036291 | 34895.91 | 10.4601
|  | 266.9482 | *K | 0.013666 | T | 274.8044 | *K | 0.007001 | T | John's Y520 | 21 | 11.7646 | 294.15 | 0.0033996 | 17382.29 | 9.7632
| Resistance | 39365.63 |  | 10.5806 | ln(R) | 27251.73 |  | 10.2129 | ln(R) |  | 28.2 | 20.19 | 301.35 | 0.0033184 | 12178.98 | 9.4075
[11 empty] | 29.4 | 20.95 | 302.55 | 0.0033052 | 11773.31 | 9.3736
21 | Temp in room temp | 11.7646 | *C |  |  | 18.9875 | *C |  |  |  | 7.6 |  | 280.75 | 0.0035619 | 27763.96 | 10.2315
|  | 284.9146 | *K |  |  | 292.1375 | *K
| Resistance | 17382.29 |  |  |  | 12812.24 |  |  |  | Beth's Y520 | 0.1 | 1.6544 | 273.25 | 0.0036597 | 27251.73 | 10.2129 | 0.00014096 | 0.00034454
[10 empty] | Beth's Y520 | 21 | 18.9875 | 294.15 | 0.0033996 | 12812.24 | 9.4582

## method from YosemiTech
From: 高文幸 <wenxing@yosemitech.com> | [5 empty] | Hi Anthony，
Date: Thursday, April 7, 2022 at 8:34 PM | [5 empty] | We do use Steinhart-Hart model to calculate temperature.
To: mike <mike@yosemitech.com>, Anthony Aufdenkampe <aaufdenkampe@limno.com> | [5 empty] | But the temperature calibration coefficients including Slope and Offset are working on resistance.
Cc: 孙贤 <sunxian@yosemitech.com> | [5 empty] | For example : At 5 degrees Celsius , read resistance value recorded as X.
Subject: Re: Help with temperature calibration of Y520 | [5 empty] | At 30 degrees Celsius , read resistance value recorded as Y .
Then K = (23462.8-8196.25)/(X-Y);
B = 23462.8-X*(23462.8-8196.25)/(X-Y).
| Temp C | Ref resistence (from table) | Measured Resistence (with K=1, B=0) |  | New Temp Calibration Coefficients
X | 5 | 23462.8 | 31310 | K: | 0.77186
Y | 30 | 8196.25 | 11531 | B: | -704.0275
NOTE: only measure after resetting temp K=1, B=0
Measured (after resetting temp K=1, B=0) based on interpolation from the online calculator below.
https://www.thinksrs.com/downloads/programs/Therm%20Calc/NTCCalibrator/NTCcalculator.htm
Then write K and B.
This method is worked for Conductivity too.
The turbidity is not recommended for calibration.
Best Regards,
Zoe