## YL43-example
Y532 pH Calibrator
v0.2.1: 2020-09-10 | [6 empty] | Calibration Standards: Known Values
Anthony Aufdenkampe <aaufdenkampe@limno.com> |  |  | Reference Standards | tap water | pH | at 25°C | 4 | pH | @ 25°C | 6.86 | pH | at 25°C | 9.18 | pH | at 25°C
|  |  |  | Not used in calibration |  |  | Calibration Standards: Measured Values
Instructions |  |  |  | pH | E (mV) | T (°C) | pH | E (mV) | T (°C) | pH | E (mV) | T (°C) | pH | E (mV) | T (°C)
Calibration Date: | 2020-09-10 | [5 empty] | 3.851 | 280.7963 | 20.5838 | 6.5008 | 124.1004 | 20.6961 | 8.8165 | -11.5469 | 20.9232
[7 empty] | 3.8502 | 280.8424 | 20.5835 | 6.542 | 121.6635 | 20.6973 | 8.8162 | -11.5315 | 20.9241
1. GET & paste sensor info from Modbus Runner software. | [6 empty] | 3.8511 | 280.7886 | 20.5844 | 6.5652 | 120.2951 | 20.6979 | 8.8142 | -11.4162 | 20.9235
SN: | YL4318052603 | [5 empty] | 3.8506 | 280.8194 | 20.5853 | 6.5808 | 119.3726 | 20.7 | 8.8157 | -11.5007 | 20.9235
Version: | hw 1.2,sw 2.0 | [5 empty] | 3.8502 | 280.8424 | 20.5856 | 6.5905 | 118.796 | 20.7003 | 8.8163 | -11.5392 | 20.9241
pH+NH4_N: | pH coefficients (previous) | [5 empty] | 3.8483 | 280.9577 | 20.5868 | 6.5997 | 118.2502 | 20.7018 | 8.8182 | -11.6468 | 20.9232
K1: | 102.9719 | mV at pH= 6.86 |  |  |  |  | 3.8421 | 281.3267 | 20.5874 | 6.6065 | 117.8504 | 20.7021 | 8.8188 | -11.6852 | 20.9241
K2: | 94.5699 | E0, mV at pH= 7, acidic |  |  |  |  | 3.8439 | 281.2191 | 20.5871 | 6.6117 | 117.543 | 20.7027 | 8.8199 | -11.7467 | 20.9238
K3: | 0.85452 | S0, acidic |  |  |  |  | 3.8452 | 281.1422 | 20.5868 | 6.617 | 117.2278 | 20.7036 | 8.8209 | -11.8082 | 20.9244
K4: | 102.9719 | mV at pH= 6.86 |  |  |  |  | 3.8425 | 281.3036 | 20.5871 | 6.6199 | 117.0586 | 20.7048 | 8.8207 | -11.7929 | 20.9244
K5: | 94.6722 | E0, mV at pH= 7, alkaline
K6: | 0.12403 | S0, alkaline | Average | #DIV/0! | #DIV/0! | #DIV/0! | 3.8475 | 281.0038 | 20.5858 | 6.5834 | 119.2158 | 20.7007 | 8.8178 | -11.6214 | 20.9238
|  |  | Std | #DIV/0! | #DIV/0! | #DIV/0! | 0.0037138 | 0.22042 | 0.0014717 | 0.037975 | 2.2464 | 0.0028619 | 0.0022915 | 0.13409 | 0.0004568
3. Fill in reference standard known values. |  |  | %CV | #DIV/0! | #DIV/0! | #DIV/0! | 0.09652% | 0.07844% | 0.007149% | 0.5768% | 1.884% | 0.01383% | 0.02599% | -1.154% | 0.002183%
Type into pale yellow cells H3, L3, P3
4."GET pH" & paste measured values for each standard. |  |  |  | Calibration Curve
Measure 3-10 "Times" to get best values. |  |  |  | X | Y |  |  |  | 2-point | [6 empty] | X | Y
Paste into pale yellow cells. |  |  |  | pH | E (mV) | T (°C) | predicted Y | residual Y | predicted X | [6 empty] | pH | E (mV) | T (°C)
|  |  |  | 4 | 281.0038 | 20.5858 | 280.9376 | 0.066279 | 4 | [6 empty] | 4 | 281.0038 | 20.5858
5. Delete calibration X:Y pairs that are not valid. |  |  |  | 6.86 | 119.2158 | 20.7007 | 119.3637 | -0.14798 | 6.86 | [6 empty] | 6.86 | 119.2158 | 20.7007
Delete green values in pale yellow cells D25:E27 |  |  |  | 9.18 | -11.6214 | 20.9238 | -11.7031 | 0.081706 | 9.18 | [6 empty] | 9.18 | -11.6214 | 20.9238
6. Re-SET "User Coefficients" to new calibration values. |  |  |  | pHiso | 7 | isopotential pH
Paste these blue values into Modbus Runner and hit "SET".
K1: | 119.2158 | mV at pH= 6.86 |  |  | S, slope | E0 | S0, slope correction
K2: | 111.2961 | E0, mV at pH= 7, acidic |  |  | mV/pH | mV at pHiso
K3: | -1.7378 | S0, acidic |  | acidic pH 4 to 6.86 | -56.5693 | 111.2961 | -1.7378
K4: | 119.2158 | mV at pH= 6.86 |  | alkaline pH 6.86 to 9.18 | -56.3953 | 111.3204 | -1.9559
K5: | 111.3204 | E0, mV at pH= 7, alkaline
K6: | -1.9559 | S0, alkaline |  | ideal slope: | -58.307 | mV/pH
7. Save this sheet for your records.
Rename tab with serial number and date, or keep a workbook for each serial number.
Background Information
Explanation of pH calibration equations, from YosemiTech
Nernst equation :E=E0+S*(7-pH) |  |  | S = 54.2+0.1984*t(℃) +S0 |  |  | Where S0 = 0,E0=0 by default.
K1=K4 = E (a sensor reading of mV at ph = 6.86)
K2 is actual E0 in acidic solution ph =4 |  |  |  |  | K5 is actual E0 in alkaline solution ph =9.18
K3 is actual S0 in acidic solution ph =4 |  |  |  |  | K6 is actual S0 in alkaline solution ph =9.18
If E1 is read at pH=4.00,E2 is read at pH = 6.86, E0 = (E2*(4.00-7)-E1*(6.86-7))/(4.00-6.86)
Bier2018-Hach-ElectrochemTheoryPractice_DOC182.53.90629.pdf
https://www.google.com/url?sa=t&rct=j&q=&esrc=s&source=web&cd=&ved=2ahUKEwiEhbGrxN_rAhUMbc0KHXdmAJ8QFjAJegQIBBAB&url=https%3A%2F%2Fwww.hach.com%2Fasset-get.download.jsa%3Fid%3D7672493965&usg=AOvVaw0aMsv1tmXVclWovq6qngYQ
Nerstt equation for pH:
| slope = 2.303*RT/nF = -0.1984*T = | -59.153 | mV/pH at 25°C | -54.193 | Adjustment for °C vs °K
| R = | 8.3145 | J K−1 mol−1
| F = | 96485.332 | C mol−1
| 2.303*R/nF = | 0.19846 | mV/K, where 2.303 converts from ln to log10; 1000 mV/V; 1 C / 1 V
Theory and Practice of pH Measurement, Chapter 3 & Appendix B
https://www.emerson.com/documents/automation/manual-theory-practice-of-ph-measurement-rosemount-en-70736.pdf
https://en.wikipedia.org/wiki/Nernst_equation